## Slide 1
Bob Dass (FAS, Mid-Atlantic)
Mahasti Alavi (Sales Manager, Mid-Atlantic)
Octet BLI Technology and Application Overview
Notes: Hello everyone
I’m going to give a brief overview of Product marketing activities for the Allegro product range
NEXT SLIDE

## Slide 2: AGENDA
AGENDA:
Octet and BLI Introduction and Over view - 30 mins
2. Popular Applications/ Assays using Octet - 10 Mins
3. Q and A from the slides presented 15 mins
4. Specific questions:
Yiran Wang, Ph.D
Buenafe T. Arachea, Ph.D.
Iga Kucharska, Ph.D

## Slide 3: ForteBio BLI Systems: Label-Free Protein Analysis
Definitive Kinetic Analysis
Analyze Unpurified Samples
Straightforward Operation
Fast Protein Quantitation

## Slide 4: What ForteBio BLI Systems Provide
Large menu of biosensors enables diverse applications
Quantitation
Direct, 1-step
Sandwich
ELISA
mg/mL – pg/mL
Kinetics
ka, kd, KD
Proteins, Abs
Peptides, oligos
Small molecules
Specificity
Function testing
Epitope binning
Rank ordering
Isotyping
Notes: What information does one gain from the Octet? The range is quite versatile – from kinetics data – information such as Ka, Kd and affinity constants for a wide range of molecules to Quantitation consisting of simple single step quantation assay up to more complex multi step assays such as Elisa (which I’ll get into more detail in a later slide). Another very valuable way the Octet can help you with your work is to gain specificity information about your molecules. By that I mean such things as functionality testing, epitope binning, ranking order, isotyping, competition assay and the list goes on.

## Slide 5: ForteBio BLI Systems: Blitz
Needs only 4 µL of sample in drop holder.
Can use 250 µL in tube holder.

## Slide 6: Data Collection on the Blitz
All steps are prompted by the software

## Slide 7: Stable Conditions Make for Stable Measurements
Do not place the instrument in direct sunlight
Allow the system at least 30 minutes to warm up after switching on the power
Bring all reagents to room temperature before analyzing

## Slide 8: High Sensitivity Measurements Require Thorough Cleaning
Insufficient cleaning of the drop holder can give false signals
Rinse with buffer at least 3x, using a kimwipe or supplied swab
Use 0.5N HCL to periodically clean the drop holder
Multiple users may want their own drop holders
Do not allow samples to dry on the drop holder

## Slide 9: Limit Sample Evaporation for Better Results
Drop Holder measurements are limited to 5 minutes
Microfuge tube measurements are limited to 15 minutes

## Slide 10: ForteBio BLI Systems: Octet
One biosensor tray
One 96-well sample/reagent plate
8 simultaneous readings
One biosensor tray
Two 96-well or 384-well sample/reagent plates
16 simultaneous readings (Octet 384)
8, 16, 32, 48, or 96 simultaneous readings (Octet HTX)
Octet 96
Octet 384/HTX

## Slide 11: HTX Breakthrough Technology: User-Selectable Read Head
96 sensor – QK384 sensitivity
32 or 48 sensors – QK384 sensitivity
8 or 16 sensors – RED96/384 sensitivity
Notes: What sets the HTX apart?
“You don’t have to choose anymore” We can do it all, high resolution and accelerated workflow
Note. Competitors: Biacore, Proteon, Wasatch, etc. focus on either high resolution kinetics/affinity or throughput. Hydra provides both.

## Slide 12
 | BLItz | Octet QKe | Octet QK384 | Octet RED96 | Octet RED384 | Octet HTX
Molecular Weight Range | >10,000 D | > 5000 D | > 5000 D | > 150 D | > 150 D | > 150 D
# Spectrometers | 1 | 1 | 2 | 8 | 16 | 16
# Channels per Read | 1 | 1 - 8 | 1 - 16 | 1 - 8 | 1 - 16 | 1 - 96
Microplate Positions | n/a | 1 | 2 | 1 | 2 | 2
Biosensor Reracking | n/a | Yes | Yes | Yes | Yes | Yes
Robot Compatible | No | No | Yes | No | Yes | Yes
Sample Vessel Formats | Drop holder or 0.5 mL tube | 96, 96 HA | 96 / 96HA 384 / 384TW | 96 | 96 / 96HA 384 / 384TW | 96 / 96HA 384 / 384TW
Minimum Sample Volume | 4 µL | 80 µL per well | 40 µL per well | 180 µL per well | 40 µL per well | 40 µL per well
Affinity range (approximate) | 1 µM to 0.1 nM | 0.1 mM to 10 pM | 0.1 mM to 10 pM | 1 mM to 10 pM | 1 mM to 10 pM | 1 mM to 10 pM
Sample Usage | Non-destructive and recoverable |  |  |  |  | 
Temperature Control | none | 4°C above ambient to 40°C |  |  |  | 
21 CFR Part 11 Compliance | Available as option for all systems |  |  |  |  | 
96 HA 96 well microplate with half-area well size
384 TW ForteBio 384 tilted-well microplate

## Slide 13: ForteBio BLI Systems: Label-Free Protein Analysis
Cells
Bacteria
Virus
IgM
IgA
IgG, IgD, IgE
Antibody Fragments
Proteins
Peptides
Nucleotides (DNA)
Small Molecules
Atoms
200 nm
1000 nm
0.1 nm
1
1,000
100,000
1,000,000
75 nm
150
MW
Size
7 nm
1 nm
Bacteria
Virus
Antibody - Antigen
Receptor - Ligand
DNA - DNA
DNA - Protein
Antibody Fragment - Antigen
Antigen – Fusion Protein
Antibody - Peptide
Multiple Antibody Pairings
Antibody - Small Molecule
Protein - Small Molecule

## Slide 14: Bio-layer Interferometry Applications
Kinetic Applications
Protein - protein interactions
Protein - small molecule interactions
Lipisome - protein/antibody binding
Bacteria – antibody binding
Virus-like particle - protein binding
DNA aptamer binding
Glycan – protein binding
GPCR-Protein binding
Quantitative Applications
Titer determination
Rapid protein IgG quantitation
Quantitation assays for ELISA replacement
Residual Protein A contamination
Protein/Antibody Quantitation
Plant protein quantitation in crude extracts
Host-cell protein contamination
Immunogenicity (low and high affinity ADA’s)
Screening Applications
Screening proteins for crystallization studies
DNA aptamer screening
Small molecule fragment screening
Secondary screening and hit validation
DNA-DNA mismatch detection
Phage binding (phage display)
Protein/peptide/small molecule inhibition
Clone selection in media
Monitoring protein expression
Bioreactor monitoring
Epitope mapping/binning
Assay Development Applications
Media development
Process development
Antibody subtyping
Antibody pair selection

## Slide 15: Biosensor Surface Chemistry
Two-dimensional binding surface
Biocompatible Matrix - minimizes non-specific binding
Uniform
Non-denaturing
|---- 600 mm ----|

## Slide 16: Biolayer Interferometry (BLI)
Optical layer reflects simple white light; second reflection from tip of biosensor, both reach detector
Analyte binding changes thickness of bio-layer, which is measured at detector

## Slide 17: BLI Biosensor Types Available
Biosensor | Application
Antibody-Specific Capture | 
• Anti-Human IgG Fc Capture (AHC) | Human IgG Fc region, kinetic analysis
• Anti-Human IgG Fc Capture (AHQ) | Human IgG Fc region, quantitation
• Anti-Mouse Fc Capture (AMC) | Mouse IgG1, 2a & 2b Fc regions, kinetic analysis
• Anti-Mouse Fc Capture (AMQ) | Mouse IgG1, 2a & 2b Fc regions, quantitation
• Anti-Human Fab-CHI (FAB) | Fab-CH1 domains of human IgG
• Protein A (ProA) | Quantitation of various species IgG
• Protein G (ProG) | Quantitation of various species IgG
• Protein L (ProL) | Quantitation of IgG via kappa light chain
Affinity Tag Capture | 
• Streptavidin (SA) | Biotinylated ligands
• High Precision Streptavidin (SAX) | Biotinylated ligands (4% CV)
• Super Streptavidin (SSA) | Biotinylated ligands (high-density surface, small molecules)
• Anti-GST (GST) | GST-tagged recombinant proteins
• Anti-Penta HIS (HIS) | HIS-tagged recombinant proteins
• Anti-Penta HIS 2nd Gen (HIS2) | HIS-tagged recombinant proteins
• Ni-NTA (NTA) | HIS-tagged recombinant proteins
Immobilization | 
• Amine Reactive 2nd Gen (AR2G) | Covalent coupling to reactive amine groups
• Aminopropylsilane (APS) | Adsorption to hydrophobic moieties

## Slide 18: Biosensor regeneration

## Slide 19: Octet can successfully measure binding in a variety of complex matrices
Sandy soil
Pollen
Diesel Fuel
Milk
Honey
Shellfish slurry
Chicken slurry
Wheat Flour slurry
Beef slurry
Crude Corn Extracts
Lecithin (Brown Fat)
Many types of media and blood plasma types
Methanol
Cell lysates
DMSO

## Slide 20: Information Provided by Octet Platform
Protein/Antibody Concentration
  Picogram to milligram dynamic range
  Contaminant testing
    Immunogenicity
    Residual Protein A Detection
    HCP Detection
Affinity characterization of proteins, antibodies, and small molecules
  Binding KD
  On and off rates
  Rank order
  Yes / No kinetic screening
  Epitope binning / mapping
Signal
Concentration

## Slide 21: Quantitation Biosensors
Biosensor | Application
Antibody-Specific Capture | 
• Anti-Human IgG Fc Capture (AHC) | Human IgG Fc region, kinetic analysis
• Anti-Human IgG Fc Capture (AHQ) | Human IgG Fc region, quantitation
• Anti-Mouse Fc Capture (AMC) | Mouse IgG1, 2a & 2b Fc regions, kinetic analysis
• Anti-Mouse Fc Capture (AMQ) | Mouse IgG1, 2a & 2b Fc regions, quantitation
• Anti-Human Fab-CHI (FAB) | Fab-CH1 domains of human IgG
• Protein A (ProA) | Quantitation of various species IgG
• Protein G (ProG) | Quantitation of various species IgG
• Protein L (ProL) | Quantitation of IgG via kappa light chain
Affinity Tag Capture | 
• Streptavidin (SA) | Biotinylated ligands
• High Precision Streptavidin (SAX) | Biotinylated ligands (4% CV loaded SA)
• Super Streptavidin (SSA | Biotinylated ligands (high-density surface)
• Anti-GST (GST) | GST-tagged recombinant proteins
• Anti-Penta HIS (HIS) | HIS-tagged recombinant proteins
• Anti-Penta HIS 2nd Gen (HIS2) | HIS-tagged recombinant proteins
• Ni-NTA (NTA) | HIS-tagged recombinant proteins
Immobilization | 
• Amine Reactive 2nd Gen (AR2G) | Covalent coupling to reactive amine groups
• Aminopropylsilane (APS) | Adsorption to hydrophobic moieties

## Slide 22: Octet Automated Workflow for Quantitation
Standards
Test Samples
Binding (nm)
Time (sec)
Binding Rate
Concentration
Data is taken for 2minutes per 8 wells
  96 wells in ~30 minutes
  1 step, no washing
Octet Biosensors
120

## Slide 23: Octet Workflow for Quantitation with Regeneration
Standards
Test Samples
Binding (nm)
Time (sec)
Binding Rate
Concentration
The binding rates of test samples are measured and interpolated from the standard curve to determine concentration
96 samples analyzed in 15 - 30 minutes
Reuse of standard curve is optional
Octet Biosensors
120
Buffer/Neut.
Regen
Notes: 23

## Slide 24: Octet Assays Successfully Replace HPLC in CCD at Biogen Idec
Biogen IDEC:
  Fc-fusion protein
  Protein G HPLC vs Octet ProG Quant
  96 samples: 19 hr reduced to 20 min
  Intermediate assay precision %CV <5% (n=54)
100%, 50%, 25%, and 12.5% CCM had comparable standard curves with overall combined CV <8%

## Slide 25: Octets can Replace Lengthy ELISAs

## Slide 26: Octet Reproducibility and Comparison to ELISA
ELISA | 5 - 20%
Octet | < 5%
Typical RSD Values:
Data presented at IBC Antibody Production, 2008 by Keith Davis from Pfizer MO.
Notes: Data from Keith Davis, Pfizer St. Louis.
The Octet does not suffer from the variability introduced to the assay by pipetting, coating, blocking, and repeated washing. The factors can also compound.

## Slide 27: Real-Time Kinetics Provides Better Binding Information Over ELISA
ELISA methods only provide approximations of KD.
Different combinations of kon and koff can give the same affinity measurement.
These two compounds appear to have the same KD, however Example B has a 100x tighter dissociation than Example A. For this pharmaceutical application, a tighter dissociation would be preferred.
Washing steps in ELISA methods will remove weak antibody interactions, and thus are not feasible to use for characterization.
 | Association | Dissociation | KD = Affinty
Example A | 1.00E+04 | 1.00E-02 | 1.00E-06
Example B | 1.00E+02 | 1.00E-04 | 1.00E-06

## Slide 28: Kinetic Biosensors
Biosensor | Application
Antibody-Specific Capture | 
• Anti-Human IgG Fc Capture (AHC) | Human IgG Fc region, kinetic analysis
• Anti-Human IgG Fc Capture (AHQ) | Human IgG Fc region, quantitation
• Anti-Mouse Fc Capture (AMC) | Mouse IgG1, 2a & 2b Fc regions, kinetic analysis
• Anti-Mouse Fc Capture (AMQ) | Mouse IgG1, 2a & 2b Fc regions, quantitation
• Anti-Human Fab-CHI (FAB) | Fab-CH1 domains of human IgG
• Protein A (ProA) | Quantitation of various species IgG
• Protein G (ProG) | Quantitation of various species IgG
• Protein L (ProL) | Quantitation of IgG via kappa light chain
Affinity Tag Capture | 
• Streptavidin (SA) | Biotinylated ligands
• High Precision Streptavidin (SAX) | Biotinylated ligands (4% CV loaded SA)
• Super Streptavidin (SSA | Biotinylated ligands (high-density surface)
• Anti-GST (GST) | GST-tagged recombinant proteins
• Anti-Penta HIS (HIS) | HIS-tagged recombinant proteins
• Anti-Penta HIS 2nd Gen (HIS2) | HIS-tagged recombinant proteins
• Ni-NTA (NTA) | HIS-tagged recombinant proteins
Immobilization | 
• Amine Reactive 2nd Gen (AR2G) | Covalent coupling to reactive amine groups
• Aminopropylsilane (APS) | Adsorption to hydrophobic moieties

## Slide 29: Automated Kinetic Characterization on the Octet
Buffer
Capture molecule
Analyte of Interest
Binding (nm)
Time
Baseline
Loading
Baseline
Association
Dissociation
8 or 16 samples can be analyzed in parallel
Measure on rates and off rates, multiple binding models
Data is displayed in real-time
Experimental protocols can be customized
Octet Biosensors

## Slide 30: Para Thyroid Hormone (PTH) Binding to PTH Receptor (Octet) Biotin-PTH1R ECD on SA biosensor – PTH(1-34)NH2 in solution. (Biacore) PTH1R on CM5 chip – PTH(1-34)NH2 in solution. Van Andel Research Institute, Massachusetts General Hospital
Instrument | KD
Biacore X | 4.9 µM
ForteBio Octet RED | 2.8 µM
Microcal iTC | 3.4 µM
PTH is of clinical interest for its ability to stimulate bone formation.
(Octet) J. Biol. Chem. 2009, 284, 28382-28391.
(Biacore) Biochemistry, 2000, 39, 8878-8887.
Notes: Overview:
Van Andel scientists performed experiments on Octet RED system, which are published in the JBC2009 paper. In the paper, they mentioned good match of results to biacore and ITC data from previously published Biochemistry 2000 paper.
Experiment details:
Biotinylated para thyroid hormone (PTH) receptor was loaded on Streptavidin biosensors and dipped into solutions of PTH protein in solution. Biacore experiments were performed on CM5 chip (amine reactive chemistry).
Conclusions:
Good match of data between Octet, Biacore and ITC as per publication by academic users. Data further shows reliability of biotinylation approach on Octet systems.

## Slide 31
APPLICATIONS Of THE OCTET –BLI TECHONOLOGY

## Slide 32: Rapidly Develop Optimal Assay Conditions
1400.0
Peptide 1: Antibody
KD = 2.6nM
Peptide 2: Antibody
KD = 460pM
1400.0
1600.0
1800.0
2000.0
1000.0
1200.0
1400.0
50mM HEPES + 0mM NaCl
+50mM NaCl
+100mM NaCl
+150mM NaCl
0% v/v Tween
0.001% v/v Tween
0.005% v/v Tween
0.01% v/v Tween
1000.0
1200.0
1400.0
Step 1:
Test two antibodies at four concentrations for affinity to target
Step 2:
Test optimal NaCl concentration
Step 3:
Test optimal Tween concentration
Step 4:
Test optimal concentration of carrier protein.
All experiments were completed
in one afternoon.
Successful conditions were used for HTP bead-based assay.
0 mg/ml BSA
0.1mg/ml BSA
0.5mg/ml BSA
1.0 mg/ml aBSA
Notes: Rapid assay development for an alpha screen assay. Best conditions from each step were carried forward into the next step. The last assay is an accumulation of all conditions. The customers normal methodology for developing these conditions would have taken 1+ weeks.

## Slide 33: Detection of Recombinant 34kD Antigen (300 nM) Binding to Specific Monoclonal Antibody in Complex Matrices
Octet RED96 parameters:
Standard 96W plate
mAb immobilized on AMC Biosensors
300C, 1000 RPM
10 minute association
Matrix and Buffer neg ctl run for each matrix tested
300nM rAg run in buffer as pos ctl
Qualitative analysis: yes/no binding of 300nM rAg in various dilutions of matrix
Sage Pollen 1:20 – 1:1280
Sandy Soil 1:20 – 1:1280
Diesel Fuel 1:200 – 1:12800
No matrix effect in dilutions of pollen and diesel fuel tested
Despite small matrix effect at more concentrated dilutions of soil, the Octet can successfully differentiate +/- binding
matrix dilutions
matrix dilutions
matrix dilutions
buffer and matrix neg ctls
buffer and matrix neg ctls
buffer and matrix neg ctls

## Slide 34: Binding of Antibody 5F12 to His-Tagged Protein on Octet RED
Binding (nm)
Time (sec)
His-tagged
Protein Loading
Wash
Association
Dissociation
10 nM
5 nM
2.5 nM
1.25nM
625 pm
313 pM
156 pM
Blank
Kinetics:
ka = 1.9e6
kd = 5.2e-5
KD = 27 pM
Octet Red parameters:
Standard 96W plate
Bio-Ni-NTA coated SA Biosensors
300C, 1000 RPM
10 minute association/10 minute dissociation

## Slide 35: DNA-Binding Protein Kinetic Analysis on the Octet QK
Binding of DNA-Binding Protein to Immobilized Biotinylated ssDNA
Rapid Analysis of Binding of Transcription Factors and Other Promoter Elements to Specific DNA Sequences
kd | ka | KD
4.78E-05 | 5.59E+04 | 8.55E-10
Biotin-
DNA
Protein
Dissociation in Buffer
Notes:   Binding buffer = 20 mM Tris, 150 mM KOAc, 10 mM MgOAc
Octet QK

## Slide 36: Detection of Lecithin (Brown Fat) binding to 6kDa peptide by OctetRED96
Octet RED96 parameters:
Standard 96W plate
6 kDa peptide coated SA Biosensors
300C, 1000 RPM
10 minute association/10 minute dissociation all phosphate buffer pH 6.5
Kinetic parameters determined using non-linear Global Fitting
Octet can successfully monitor interaction of very sticky, fatty molecules to accurately measure kinetic parameters.

## Slide 37: Measuring Avidity on the OctetRED96
S1, KD=1.81nM
S2, KD=0.84nM
S3, KD=2.34 nM
S4, KD=0.81nM
S5, KD=0.37 nM
Octet Red96 parameters:
Standard 96W plate
Antigen coated SA Biosensors
300C, 1000 RPM
5 minute association/10 minute dissociation all in in 4%, 2% or 1% serum
Kinetic parameters determined using non-linear Global Fitting
Real-time results available immediately allowing fast-
ranking of avidity samples.

## Slide 38: Octet Approach For Cooperative/Allosteric Binding Events
The classical allosteric interaction is: A and B binds, addition of C increases (or decreases) the binding of B to A. So one needs to figure out how C affects the binding. In many cases the allosteric effects occur due to the change of conformation of B upon the addition of C, thus affecting the interaction between A and B. Receptor is typically undergoes such conformational change.
Approach: (1) Determine pair-wise affinity (interaction), A - B, B - C, C - A. (2) if we assume C changes the conformation of B, perform A-B interactions (fixed A and B conc, say A is immobilized and B is in solution), with increased concentrations of C into analyte, to titrate for C concentration needed to see the new A-B binding profile until the A-B interaction no longer changes; (3) Use this C concentration throughout in the matrix, keep A immobilized, titrate B concentrations to get full kinetics of A and B under the cooperative conditions.

## Slide 39: Publications on Cooperative/Allosteric Binding Using BLI as Detection Platform
An Unusual Repressor Controls the Expression of the Crucial Nicotine-Degrading Gene Cluster in Pseudomonas Putida S16; Molecular Microbiology (2014)
Biolayer Interferometry for Measuring Kinetics of Protein-Protein Interactions and Allosteric Ligand Effects; Journal of Visualized Experiments ((2014)

## Slide 40: Cooperative Binding of 2 Nic2 dimers to target DNA
Binding signals at equilibrium evaluated. Hill plot for NicR2 binding to fragment b shown below.
Experiments performed on Octet RED96
An Unusual Repressor Controls the Expression of a Cruicial Nicotine-Degrading Gene Cluster in Pseudomonas putida S16, Molecular Microbiology (2014)

## Slide 41
Hit Identification and Lead Confirmation in the Discovery of Drugs Targeting Bromodomain Proteins
Liu Liu1, Yujun Zhao1, Xu Ran1, Yongqiang Zhu1, Longchuan Bai1, Donna McEachern1, Chao-Yie Yang1, Jennifer Meagher2, Jeanne Stuckey2, Shaomeng Wang1
Department of Internal Medicine and Comprehensive Cancer Center,
2. Life Sciences Institute,
University of Michigan, Ann Arbor, MI 48109

## Slide 42
Application of the Octet RED96 system
Assay development and optimization after target determined
accuracy: compared with results from other well- established assays; misleading data is worse than no data
reproducibility, stability, cost, feasibility
Find proper sensors: SSA sensors for small molecules; possible less expensive alternatives
Protein immobilization conditions: Biotinylation, protein concentrations, buffer conditions
Assay condition optimization, data interpretation
Help needed from the manufacturers, particularly in the initial stages; fairly quick and straightforward

## Slide 43
KD values of reference compounds to BRD4BD2 determined by BLI
KD = 100nM
KD = 380nM
KD = 54nM
CMPDs |  | IBET-762 | PFI-1
MW |  | 424 | 374
BRD4BD1 | OctetRED | 147 | 131
 | Published | 53* | 136
BRD4BD2 | OctetRED | 54 | 380
 | Published | 53* | 303
* BRD4 BD1-2, both reported data and our own data suggested no significant difference between BD1-2 and BD1 and BD2.
Assays developed in a couple days
Super Streptavidin (SSA) sensors, PBS with 0.1% BSA, 0.01% Tween-20, 0.5% DMSO
Inactive blocked Streptavidin and buffer as double references

## Slide 44
Screening to BRD4BD2
250µM, 0.5% DMSO, finished in a few hours
Positive hits: > average of buffer + 3 SD
Hit 1, 2, 3 are derivatives from the same scaffold.
Hit 1
Hit 2
Hit 3

## Slide 45
Hit confirmation by dose dependence titration
CMPD | MW | KD (µM) | kon (1/Ms) | kdis (1/s)
Hit 1 | 133 | 89.2 | 2.21E+02 | 1.97E-02
Hit 2 | 210 | 42.5 | 5.57E+02 | 2.37E-02
Hit 3 | 211 | 14.9 | 1.52E+03 | 2.27E-02

## Slide 46
Binding affinities of Hit 3 to other BET-BRDs
BETs | KD (µM)
BRD3BD1 | 58.7
BRD4BD1 | 14.9
BRD4BD2 | 18.3

## Slide 47: Direct Comparison of Octet RED to Biacore Published Results
Compound | ForteBio | Biacore 1,2
Acetazolamide 222 Daltons | KD = 0.039 µM | KD = 0.039, 0.019 µM
Benzenesulfonamide 157 Daltons | KD = 2.4 µM | KD = 0.8, 0.85 µM
Furosemide 330 Daltons | KD = 1.2 µM | KD = 1.0, 0.51 µM
Sulpiride 341 Daltons | KD = 239 µM | KD = 48, 186 µM
Results were within 3x and considered equivalent for affinity.
1. Papalia et al, Analytical Biochem. 359 (2006), 94-105
2. Myszka, Analytical Biochem. 329 (2004), 316-323

## Slide 48: Comparable to Biacore in a Roche Blind Study: 12/12 Match
MW (Dalton) | ForteBio (uM) | Biacore (uM)
258 | 0.560 | 0.396
162 | 0.139 | 0.458
268 | 1.24 | 3.31
282 | 4.0 | 5.7
296 | 0.635 | 0.246
331 | 10.31 | 5.57
295 | 1.46 | 1.28
269 | 11.78 | 4.63
269 | 0.261 | 0.418
221 | 1.68 | 1.88
324 | No Binding | No Binding
300 | No Binding | No Binding
Customer Requirement:
Results must be within 2-3X of Biacore values
= Equivalent
Test for non-specific
binding property
of sensors

## Slide 49: Typical Epitope Binning Assay Formats Cross-Blocking Assays To Establish Antibody Diversity
All Octet systems can perform binning in all three binning formats
Discrete, parallel and independent processing of multiple interactions (8-96 channels) simultaneously on Octet systems

## Slide 50: Octet QK enables fast Epitope Mapping / Epitope Binning
ALDER Antibody | Epitope Bin
Ab 1 | IV
Ab 2 | I
Ab 3 | I
Ab 4 | II
Ab 5 | III
Ab 6 | II
Ab 7 | II
Ab 8 | I
Ab 9 | IV
Ab 10 | IV
Easily resolved blocking of extra epitope on dimeric antigen by incubating with non-biotinylated Control Ab prior to screening other Abs
Notes: Easily resolved blocking of extra epitope on dimer antigen by incubating with unbiotinylated Control Ab

## Slide 51: Take the Pain Out of Epitope Binning
6-8 hours for analysis of matrix data
Tedious work in Excel
Not all formats work
Still being done with ELISA
Save & Export Raw Data
Microsoft Excel Macro:
Chops the long binding cycle with 16 regenerations into individual cycles & saves data as text file
Scrubber 2.0c:
Open the saved text file
Measure Binding Responses:
Antigen captured
mAb-1 binding
mAb-2 binding

## Slide 52: Biopharm Portal at www.fortebio.com

## Slide 53: User Specific Questions
Yiran Wang -
My research topic is about interaction between RNA and virus like particle. I have already tried it on BLItz and found some interaction between my virus like particle and yeast's total RNA. Now I want to get more kinetic information with BLItz or Octet. My problem is that the results are not consistent with each other even in the exact same condition. I want to use Octet to set up a suitable buffer system to continue my study.
I would like to hear more about virus like particle binding and how to adjust buffer to obtain a reliable results. Besides, I want to learn more about different models the software used.

## Slide 54
2. Buenafe T. Arachea, Ph.D. Enzymatic assays in detergent micelle, membrane lipid and nanodiscs environment examples-
3. Iga Kucharska- Interested in protein-small molecule kinetics examples

## Slide 55: Accommodating Bivalent Analytes with a Capture-Based Assay Format
Y
Y
Target molecule
Bivalent molecule
Anti Human capture sensor surface
Sensor surface
Target molecule
Bivalent molecule
Y
Y
Complex Curves
1:1 Homogeneous Curves

## Slide 56: Bivalent Analyte (Antigen-Antibody)
An artifact of the surface interaction resulting in slower dissociation (avidity).
A
+
B
A
B
k
a1
d1
k
AB
+
B
A
B2
k
a2
d2
k
Formation of AB complex
Formation of AB2 complex
Notes: This is an example of a ‘linked’ interaction, where the formation of AB2 cannot proceed before the formation of AB, and AB cannot dissociate before the dissociation of AB2

## Slide 57: Bivalent Analyte (Antigen-Antibody)
Avidity can be eliminated by optimizing (lowering) the loading concentration
Therefore even though you have a bivalent molecule, you have optimized for 1:1 binding
A
+
B
A
B
k
a1
d1
k
Formation of AB complex
Formation of AB complex
Notes: This is an example of a ‘linked’ interaction, where the formation of AB2 cannot proceed before the formation of AB, and AB cannot dissociate before the dissociation of AB2

## Slide 59: Kinetic Characterization of Protein:Protein Interactions
Example:
3) You can identify which clones have well defined binding interactions and which have complex binding interactions
Well behaved 1:1 binding interaction
Undesirable complex binding interactions
Binding saturates
Single exponential
Returns to baseline
Super-stoichiometric binding
Biphasic on, off curves
No NSB!
Specific binding
+ non-specific binding?

## Slide 60: Heterogenous Binding
Heterogeneous binding signal is typically the result of a non-specific binding interaction contaminating the specific binding signal
Take steps to eliminate the non-specific interaction
  Analyte concentrations may be too high
  Ligand loading level may be too high
  Association time might be too long
  Assay buffer may need some optimization to reduce non-specific interactions
  Reference sensor may show similar non-specific signal increase, which may be subtracted from the binding signals

## Slide 61: Different Types of NSB
Bare Sensor
Analyte
Loaded Sensor
Analyte
NSB
NSB

## Slide 62: Minimizing NSB on Biosensors
Bare Sensor
Analyte
Buffer Optimization
  BSA
  Salt
  Detergent
Biotinylated irrelevant protein to block surface
Biotinylated BSA to block surface
Biocytin to block SA binding pockets
More BSA to block surface (mg/mL range)
Other options for blocking agents:
  Non-fat milk
  Casein
  Irrelevant IgG (mg/mL range)

## Slide 63: NSB Mechanisms
Coulombic interaction
  Charge of surface, charge of molecule
  Isoelectric Point (pI) of analyte
  Ionic strength of solvent is low
Non-covalent binding
  Van der Walls interactions
  Hydrophobic interactions
Analyte solubility
  Organic compounds
  Aqueous solvent insolubility
Others
  Biological similarities
  Streptavidin to Fibronectin
Buffer pH/salt
concentration
Raise surfactant level in buffers
Add BSA and/or Casein
Raise DMSO content in assay buffer
Reduce concentration of compound
Add glycerol
Blocking open binding sites
with Biocytin

## Slide 64
THANK YOU
Mahasti Alavi, Mid-Atlantic Sales Manger
Mahasti_alavi@pall.com
Bob Dass, Mid-Atlantic FAS
Bob_dass@pall.com

## Slide 65
Heterogeneous Ligand Model
The 2:1 heterogeneous ligand model assumes analyte binding at
two independent ligand sites. Each ligand site binds the analyte
independently and with a different rate constant. Two sets of rate
constants are given, one for each interaction:
A + B1 AB1
ka1
kd1
A + B2 AB2
ka2
kd2
where A represents the analyte and B represents the immobilized
ligand.

## Slide 66
1:2 Bivalent Analyte
The 1:2 Bivalent Analyte model fits the binding of one bivalent
analyte to a monomeric immobilized ligand. Kinetic parameters are
calculated for two interactions (ka1, ka2, kd1, kd2, KD1, KD2).
A+B AB AB+B AB2
ka1
kd1
ka2
kd2
This model assumes that because of limited distance between two
adjacent binding sites on the surface, the bivalent analyte can form
a bridged complex.

## Slide 67: Assay Development

## Slide 68: Microplates
A minimum of 2 microplates is required for every Octet assay (pre-wet plate and sample plate)
Octet QKe and Octet Red96: 96 well plates
Octet QK384 and Octet Red384 and HTX can use 96 well plate or 384 well plates
96 well plates: 200uL volume required
384 well plates: 120uL volume required
384 tilted-well plates (Red384 only): 50uL
Plates need to be Greiner black, polypropylene, flat bottom plates
  Black 96 well plates: Greiner catalogue # 655209
  Black 384 well plates: Greiner catalogue # 781209
  Black 384 tilted-well plates: ForteBio 18-5080 or 18-5076
Sample is recoverable and measurement non-destructive.

## Slide 69: Part 2: Quantitation on the Octet

## Slide 70: Quantitation Biosensors
Biosensor | Application
Antibody-Specific Capture | 
• Anti-Human IgG Fc Capture (AHC) | Human IgG Fc region, kinetic analysis
• Anti-Human IgG Fc Capture (AHQ) | Human IgG Fc region, quantitation
• Anti-Mouse Fc Capture (AMC) | Mouse IgG1, 2a & 2b Fc regions, kinetic analysis
• Anti-Mouse Fc Capture (AMQ) | Mouse IgG1, 2a & 2b Fc regions, quantitation
• Anti-Human Fab-CHI (FAB) | Fab-CH1 domains of human IgG
• Protein A (ProA) | Quantitation of various species IgG
• Protein G (ProG) | Quantitation of various species IgG
• Protein L (ProL) | Quantitation of IgG via kappa light chain
Affinity Tag Capture | 
• Streptavidin (SA) | Biotinylated ligands
• High Precision Streptavidin (SAX) | Biotinylated ligands (4% CV loaded SA)
• Super Streptavidin (SSA | Biotinylated ligands (high-density surface)
• Anti-GST (GST) | GST-tagged recombinant proteins
• Anti-Penta HIS (HIS) | HIS-tagged recombinant proteins
• Anti-Penta HIS 2nd Gen (HIS2) | HIS-tagged recombinant proteins
• Ni-NTA (NTA) | HIS-tagged recombinant proteins
Immobilization | 
• Amine Reactive 2nd Gen (AR2G) | Covalent coupling to reactive amine groups
• Aminopropylsilane (APS) | Adsorption to hydrophobic moieties

## Slide 71: Octet Automated Workflow for Quantitation
Protein A sensor
IgG standard
Purified sample,
Cell culture supernatant,
Crude lysate, etc.
Capture Protein
Notes: How does it work? This is the tip of a fiber. We coat the tip end with a special optical layer. The capturing molecule is then attached to the tip. The tip is dipped into the sample containing target molecule. The target molecule binds to the capture molecule, and two form a molecular layer. A white light is directed into the fiber. Two beams will be reflected to the back end. The first beam comes from the tip as a reference. The second light comes from the molecular layer. The difference of two beams will cause a spectrum color pattern as shown here. The phase is a function of the molecular layer thickness and corresponding to the number of molecules on the tip surface.

## Slide 72: Octet Workflow for Quantitation
Standards
Test Samples
Binding (nm)
Time (sec)
Binding Rate
Concentration
The binding rates of test samples are measured and interpolated from the standard curve to determine concentration
96 samples analyzed in 15 - 30 minutes
Reuse of standard curve is optional
Octet Biosensors
120
Notes: 72

## Slide 73: Octet Workflow for Quantitation with Regeneration
Standards
Test Samples
Binding (nm)
Time (sec)
Binding Rate
Concentration
The binding rates of test samples are measured and interpolated from the standard curve to determine concentration
96 samples analyzed in 15 - 30 minutes
Reuse of standard curve is optional
Octet Biosensors
120
Buffer/Neut.
Regen
Notes: 73

## Slide 74: Instrument Flexibility Offers Wide Dynamic Range Octet RED96/RED384
mg/ml …Dynamic range...pg/mL
Y
Y
Y
Y
Y
Y
1: Analyte
HRP-tagged antibody
Precipitating substrate
Y
Y
Y
Y
Y
Y
Y
Y
2: Sandwich
1: Analyte
3: Amplification
Results will vary with assay and matrix
Direct 400 = 2000 -0.5 µg/mL
Flow Rate:
Sample Incubation:
Amplification Step(s):
400RPM
2 min/column
N/A
1000RPM
30-60 minutes
5 min/column
Direct 1000 = 300 -0.025 µg/mL
Sandwich = 500 -0.1 ng/mL
Direct 400
Y
Y
Y
Y
Y
Y
1: Analyte
1000RPM
5 min/column
N/A
Direct 1000
Y
Y
Y
Y
Y
Y
Y
Y
Y
2: Sandwich
1: Analyte
1000RPM
30-60 minutes
2 min/column
Sandwich
ELISA
ELISA = 200 -0.01 ng/mL

## Slide 75: Time and Shake Speed Optimization
Expected conc (µg/mL) | 200 rpm 2 min read time |  | 1000 rpm 5 min read time | 
 | Ave conc N=5 (µg/mL) | %CV N=5 | Ave conc N=4 (µg/mL) | %CV N=4
2000 | 2157.3 | 9% | --- | ---
1500 | 1481.6 | 3% | --- | ---
1000 | 958.2 | 2% | --- | ---
700 | 696.9 | 2% | --- | ---
500 | 485.5 | 1% | --- | ---
300 | 308.3 | 1% | 301.0 | 2%
100 | 105.6 | 1% | 100.1 | 1%
30 | 29.9 | 1% | 30.1 | 2%
10 | 9.3 | 2% | 10.0 | 1%
3 | 3.2 | 5% | 3.0 | 5%
1 | 1.0 | 2% | 1.0 | 1%
0.50 | 0.5 | 3% | 0.50 | 2%
0.25 | --- | --- | 0.25 | 3%
0.10 | --- | --- | 0.10 | 4%
0.05 | --- | --- | 0.05 | 8%
200 RPM
1000 RPM
All data on Octet RED

## Slide 76: Part 3: Kinetic Characterization on the Octet

## Slide 77: Kinetic Biosensors
Biosensor | Application
Antibody-Specific Capture | 
• Anti-Human IgG Fc Capture (AHC) | Human IgG Fc region, kinetic analysis
• Anti-Human IgG Fc Capture (AHQ) | Human IgG Fc region, quantitation
• Anti-Mouse Fc Capture (AMC) | Mouse IgG1, 2a & 2b Fc regions, kinetic analysis
• Anti-Mouse Fc Capture (AMQ) | Mouse IgG1, 2a & 2b Fc regions, quantitation
• Anti-Human Fab-CHI (FAB) | Fab-CH1 domains of human IgG
• Protein A (ProA) | Quantitation of various species IgG
• Protein G (ProG) | Quantitation of various species IgG
• Protein L (ProL) | Quantitation of IgG via kappa light chain
Affinity Tag Capture | 
• Streptavidin (SA) | Biotinylated ligands
• High Precision Streptavidin (SAX) | Biotinylated ligands (4% CV loaded SA)
• Super Streptavidin (SSA | Biotinylated ligands (high-density surface)
• Anti-GST (GST) | GST-tagged recombinant proteins
• Anti-Penta HIS (HIS) | HIS-tagged recombinant proteins
• Anti-Penta HIS 2nd Gen (HIS2) | HIS-tagged recombinant proteins
• Ni-NTA (NTA) | HIS-tagged recombinant proteins
Immobilization | 
• Amine Reactive 2nd Gen (AR2G) | Covalent coupling to reactive amine groups
• Aminopropylsilane (APS) | Adsorption to hydrophobic moieties

## Slide 78
Buffer
Ligand-Biotin
Protein of Interest
Binding (nm)
Time
Baseline
Loading
Baseline
Association
Dissociation
8 or 16 samples can be analyzed in parallel
Measure on rates and off rates
Data is displayed in real-time
Experimental protocols can be customized
Baseline, Association, and Dissociation steps on consecutive order are necessary for full KD analysis
Octet Biosensors
Octet Automated Workflow for Kinetics
Notes: Using strepavidin coated biosensors:
Step 1 is Baseline- buffer only- want nice flat line, showing zero NSB
Step 2 load B-Ab- want gradual increase that eventually flattens out.
Step 3 is second basline to rid away any excess B-Ab
Step 4 is association step of your protein of interest
Step 5 is measure the dissociation of your protein of interest

## Slide 79: Setting up Kinetics Assays
Set Shake speed to 1000rpm.
Remember: the analysis software needs Baseline, Association and Dissociation steps consecutively

## Slide 80: Using the KD Value to Develop Assays
Use assay concentrations between 0.1x KD to 10x KD
For unknowns, use sample neat and titrate down
%age of Rmax
Time
100
50
Saturation (Rmax) at 100x KD
Equilibration at 50% at KD
Detection limit around 0.1x KD
10x KD

## Slide 81: QUESTION FOR THE AUDIENCE
Q: What’s wrong with these data sets?
TO LOW!
TOO HIGH!
Just right!
A: The analyte concentration series has not been optimized for an ideal kinetic analysis